Detailed Technical Documentation

Improved Korotkoff Duration Detection Analysis v3.0

1. Introduction

This project presents an advanced RF-based Korotkoff sound detection and validation framework using:

- RF radar sensing

- Stethoscope audio validation

- Signal processing

- Statistical detection algorithms

- Multi-method consensus analysis

- Confidence estimation

- Biomedical feature extraction

The primary goal of the system is to detect Korotkoff sound onset, offset, duration, and heartbeat synchronization from RF displacement signals and compare the results with conventional stethoscope measurements.

The framework combines:

- adaptive thresholding,

- change-point detection,

- time-frequency analysis,

- heartbeat synchronization,

- statistical confidence intervals,

- and multimodal cross-validation.

2. Overall System Workflow

The entire workflow of the proposed framework is:

- RF IQ data acquisition

- IQ imbalance correction

- Phase extraction and unwrapping

- RF motion displacement estimation

- Bandpass filtering

- Korotkoff velocity extraction

- Multi-resolution energy computation

- CUSUM change-point detection

- Six-method consensus analysis

- Per-beat Korotkoff validation

- Bootstrap confidence interval estimation

- Stethoscope cross-validation

- Confidence scoring

- Dashboard visualization

3. RF Signal Processing Pipeline

3.1 RF IQ Data Loading

The system loads radar IQ data from:

with h5py.File(rf_path, 'r') as f:

data = f['data'][:]

The RF system operates at:

- Carrier frequency = 0.9 GHz

- Sampling frequency = 10 kHz

The I/Q channels represent:

- In-phase component (I)

- Quadrature component (Q)

These signals contain micro-motion information related to:

- vessel vibration,

- heartbeat motion,

- and Korotkoff turbulence.

4. IQ Conditioning and Phase Extraction

4.1 IQ Mapping

The IQ signal is constructed using:

-i + 1j * q

This converts the RF channels into a complex-valued signal.

4.2 IQ Imbalance Correction

Function:

iq_condition()

Purpose:

- removes phase imbalance,

- amplitude mismatch,

- quadrature distortion.

The algorithm uses:

- cross-correlation orthogonalization,

- covariance normalization,

- adaptive correction.

This improves:

- phase stability,

- displacement estimation,

- signal-to-noise ratio.

4.3 Robust Phase Extraction

Function:

robust_phase()

Steps:

- Apply 50 Hz low-pass filter

- Compute phase angle

- Perform phase unwrapping

- Remove phase drift using detrending

Mathematically:

Phase = unwrap(angle(IQ))

This phase contains:

- chest displacement,

- vascular motion,

- Korotkoff vibration information.

5. RF Biomedical Feature Extraction

5.1 Korotkoff Velocity Extraction

Bandpass filter:

10–49 Hz

Reason:

Korotkoff vibrations exist primarily in this frequency region.

The filtered phase derivative gives:

Velocity = d(phase)/dt

The system converts phase into physical displacement using:

Scale = λ / (4π)

where:

- λ = RF wavelength.

5.2 Heart Rate Extraction

Bandpass:

0.5–3 Hz

Purpose:

- heartbeat detection,

- pulse synchronization,

- physiological validation.

6. Multi-Resolution Energy Analysis

Function:

multi_resolution_energy()

The system computes energy using three windows:

| Window | Purpose |
| --- | --- |
| 200 ms | transient Korotkoff detection |
| 500 ms | stable pulse detection |
| 1500 ms | long-duration stability |

Each energy curve is normalized.

Then geometric mean fusion is applied:

Consensus = (E200 × E500 × E1500)^(1/3)

Advantages:

- suppresses noise,

- reduces false positives,

- improves robustness.

7. Adaptive Thresholding

Function:

adaptive_threshold()

Traditional methods use fixed thresholds.

This framework estimates:

- noise median,

- median absolute deviation (MAD),

- adaptive statistical threshold.

Threshold equation:

Threshold = median + k × MAD

Advantages:

- robust against recording variation,

- subject-independent,

- adaptive to environmental noise.

8. CUSUM Change-Point Detection

Function:

cusum_change_points()

CUSUM detects:

- sudden energy increase,

- sudden energy decrease.

These correspond to:

- Korotkoff onset,

- Korotkoff offset.

The algorithm computes:

S+ = positive cumulative sum

S− = negative cumulative sum

Advantages:

- precise onset localization,

- robust offset detection,

- better than static thresholding.

9. Six-Method Consensus Detection

The framework combines six independent methods:

| Method | Description |
| --- | --- |
| M1 | RMS Energy |
| M2 | Teager-Kaiser Energy Operator |
| M3 | Kurtosis |
| M4 | Hilbert Envelope |
| M5 | Band Power Ratio |
| M6 | STFT Energy |

Each method independently estimates:

- onset,

- offset,

- duration.

Consensus is obtained using:

Median fusion

Advantages:

- reduces detector bias,

- improves robustness,

- improves reproducibility.

10. Teager-Kaiser Energy Operator (TKEO)

Function:

calc_tkeo()

Equation:

Ψ[x(n)] = x(n)^2 − x(n−1)x(n+1)

Purpose:

- emphasizes transient oscillations,

- enhances Korotkoff pulses,

- improves weak-signal detection.

11. STFT-Based Spectral Analysis

Function:

stft()

Purpose:

- time-frequency representation,

- spectral energy estimation,

- transient sound analysis.

The spectrogram helps visualize:

- Korotkoff energy evolution,

- frequency concentration,

- onset and offset regions.

12. Per-Beat Korotkoff Detection

Function:

detect_per_beat_korotkoff()

The framework:

- Detects heartbeat peaks

- Creates beat-centered windows

- Computes Korotkoff energy per beat

- Determines active Korotkoff beats

Advantages:

- physiologically meaningful,

- heartbeat synchronized,

- clinically interpretable.

This is one of the strongest contributions of the framework.

13. Bootstrap Confidence Interval Estimation

Function:

bootstrap_onset_offset()

Purpose:

- statistical uncertainty estimation,

- robustness analysis,

- confidence interval generation.

The algorithm:

- Adds random noise

- Re-runs detection

- Re-estimates onset and offset

- Computes 95% confidence intervals

Advantages:

- statistical rigor,

- uncertainty quantification,

- publication-quality analysis.

14. Heart Rate Estimation

The framework estimates heart rate using:

14.1 Peak Detection

Using:

find_peaks()

14.2 PSD Estimation

Using:

Welch PSD

Advantages:

- dual validation,

- robust heart rate estimation,

- physiological consistency checking.

15. Stethoscope Signal Processing

The framework loads:

- MP4 audio,

- WAV audio.

The stethoscope signal is filtered into:

| Band | Purpose |
| --- | --- |
| 20–200 Hz | Korotkoff sound |
| 0.5–5 Hz | Heart rate |

This acts as:

Ground truth reference

for RF validation.

16. Cross-Validation Framework

The framework computes:

16.1 IoU (Intersection over Union)

Measures overlap between:

- RF detection,

- stethoscope detection.

16.2 Cross-Correlation

Measures:

- signal similarity,

- temporal alignment.

16.3 Lag Correction

Compensates for:

- acquisition delay,

- processing offset,

- synchronization mismatch.

17. Confidence Score System

Function:

compute_confidence_score()

The final confidence combines:

| Parameter | Weight |
| --- | --- |
| IoU overlap | 0.30 |
| Onset agreement | 0.25 |
| HR agreement | 0.25 |
| Method consensus | 0.20 |

Final output:

HIGH

MODERATE

LOW

VERY LOW

Advantages:

- interpretable validation,

- quantitative reliability,

- clinical usability.

18. Dashboard Visualization

The framework generates a:

16-panel dashboard

including:

- RF velocity,

- STFT spectrogram,

- energy curves,

- heartbeat timeline,

- per-beat Korotkoff activity,

- cross-correlation,

- confidence metrics,

- statistical summaries.

This provides:

- publication-quality figures,

- debugging support,

- clinical interpretation.

19. Major Strengths of the Framework

Strong Contributions

- Multi-resolution energy fusion

- RF + stethoscope multimodal validation

- Per-beat physiological synchronization

- CUSUM onset/offset detection

- Bootstrap confidence intervals

- Six-method consensus analysis

- Confidence scoring framework

- Publication-quality visualization

20. Current Limitations

Although the framework is highly advanced, some limitations remain:

- No cuff pressure synchronization

- Limited motion artifact suppression

- Fixed Korotkoff frequency range

- No deep learning model

- Offline processing only

- No large multi-subject validation

21. Recommended Future Improvements

21.1 Wavelet Transform

Replace STFT with:

Continuous Wavelet Transform (CWT)

Advantages:

- better transient localization,

- adaptive time-frequency resolution.

21.2 Deep Learning

Possible models:

- CNN

- Transformer

- BiLSTM

- Temporal Convolutional Networks

Input:

- RF spectrogram,

- wavelet scalogram,

- heartbeat envelope.

21.3 Motion Artifact Removal

Add:

- ICA,

- PCA,

- EMD,

- VMD.

Purpose:

- movement suppression,

- respiration removal,

- environmental noise reduction.

21.4 Pressure Sensor Fusion

Integrate cuff pressure sensor for:

- systolic estimation,

- diastolic estimation,

- complete BP measurement.

21.5 FPGA Real-Time Deployment

Convert offline framework into:

real-time embedded implementation

using:

- FPGA DSP blocks,

- streaming processing,

- hardware acceleration.

22. Research Significance

This framework demonstrates a highly advanced approach for:

- RF-based biomedical sensing,

- contactless cardiovascular monitoring,

- intelligent Korotkoff detection,

- multimodal physiological validation.

The system has strong potential for:

- publication,

- patenting,

- real-time healthcare systems,

- wearable sensing,

- AI-assisted blood pressure monitoring.

23. Conclusion

The proposed Improved Korotkoff Detection Framework v3.0 combines:

- RF sensing,

- advanced signal processing,

- statistical analysis,

- multimodal validation,

- and confidence estimation

into a highly robust biomedical sensing system.

Compared to traditional methods, the framework provides:

- improved robustness,

- adaptive detection,

- statistical reliability,

- physiological synchronization,

- and explainable validation.

The proposed architecture demonstrates strong potential for future:

- cuffless blood pressure systems,

- intelligent healthcare sensing,

- RF biomedical monitoring,

- and embedded AI healthcare applications.